To manage a VXI11 connection to your Siglent SDS1202X-E, you do not need to install massive VISA drivers (like National Instruments NI-VISA). Instead, you can use the pure Python python-vxi11 library, which is lightweight, platform-independent, and handles LXI/VXI11 communication natively over your local network. [1, 2, 3]

1. Requirements Installation

First, open your terminal or command prompt and install the native VXI11 library:

pip install python-vxi11

2. Python VXI11 Connection Template

This clean template encapsulates the connection inside a Python context manager (with block) to ensure that network sockets are safely closed, even if your script crashes. [4]

import time

import vxi11

# --- CONFIGURATION ---

# Replace with the actual static IP address displayed on your Siglent SDS1202X-E

SCOPE_IP = "192.168.1.50"

def main():

print(f"Connecting to Siglent SDS1202X-E at {SCOPE_IP} via VXI11...")

try:

# Initialize the instrument connection

with vxi11.Instrument(SCOPE_IP) as scope:

# 1. Identify the device to confirm successful communication

# Siglent devices require an explicit delay or timeout adjustments occasionally

scope.timeout = 5000  # Set timeout to 5 seconds

idn_response = scope.ask("*IDN?")

print(f"\n[SUCCESS] Connected to: {idn_response}")

# 2. Example 1: Emulate the Math button using front panel simulation

print("\nSending front panel emulation command: Press 'Math' menu...")

scope.write("SY_FP 31,1")

time.sleep(1) # Allow UI frame buffer to update

# 3. Example 2: Preferred Native SCPI trace queries

# Query the current vertical scale of Channel 1

c1_scale = scope.ask("C1:VDIV?")

print(f"Channel 1 Vertical Scale: {c1_scale.strip()}")

# Query the current horizontal timebase

timebase = scope.ask("TIME_DIV?")

print(f"Horizontal Timebase Scale: {timebase.strip()}")

# 4. Clear any errors or pending states on the instrument buffer

scope.write("*CLS")

print("\nSession completed. Connection closed automatically.")

except vxi11.vxi11.Vxi11Exception as e:

print(f"\n[ERROR] VXI11 Communication failed: {e}")

print("Please verify the oscilloscope is turned on, connected to LAN, and the IP is correct.")

except Exception as e:

print(f"\n[ERROR] An unexpected error occurred: {e}")

if __name__ == "__main__":

main()

Essential Siglent Programming Tips

- Command Termination: The python-vxi11 package handles raw line endings (\n) for you seamlessly when using .write() and .ask(). Do not explicitly append \n to your commands unless you notice truncation. [5]

- Avoid Socket Overload: Older Siglent firmware versions can occasionally drop standard RPC termination tags if queries are flooded too fast. Always include a tiny delay (time.sleep(0.05)) between continuous loops of scope.ask() commands. [4]

- Native Strings: Data returned via .ask() comes as a standard string. For extracting large, raw data arrays (like binary waveform preambles), you should transition to using scope.read_raw() after sending a write command. [5]

If you want to pull data or export specific measurements, tell me if you need help with:

- The syntax to download a full screenshot (.bmp/.png) directly to your PC.

- The SCPI sequence needed to extract raw waveform data points into a CSV or Python list.

[6, 7, 8, 9]

[1] https://siglentna.com

[2] https://www.eevblog.com

[3] https://int.siglent.com

[4] https://github.com

[5] https://www.siglenteu.com

[6] https://siglentna.com

[7] https://siglentna.com

[8] https://forums.ni.com

[9] https://pypi.org